## Test Results
Test Results of 2006 Sample Meters - Year 15
Serial No. | Maximum | Intermediate | Minimum | Total Water Flow | Testing Date
32579088 | 99.2 | 100 | 98 | 313700 | 2021-12-03
32579085 | 98.8 | 101 | 97 | 623367 | 2021-12-03
32579090 | 99 | 101 | 97 | 435464.2 | 2021-12-03

## Maximum
ANSI Standard for Maximum Flow
Test Results
1 | Sample Size: n | 3 | 99.2
2 | Sum of Measurements | 297 | 98.8
3 | Sum of Squared Measurements | 29403.08 | 99
4 | Correction Factor (CF) | 29403
5 | Corrected Sum of Squares (SS) | 0.08
6 | Variance (V) | 0.04
7 | Estimate of Lot Standard Deviation | 0.2
8 | Sample Mean | 99
9 | Upper Specification Limit | 101.5
10 | Lower Specification Limit | 98.5
11 | Quality Index: QU (upper) | 12.5
12 | Quality Index: QL (lower) | 2.5
ANSI Standard Table B-5 used to derive values below
13 | Est. of Lot Percent NcF above Upper | 0%
14 | Est. of Lot Percent NcF below Lower | 0%
15 | Total Est. Percent NcF in Lot (p) | 0%
16 | Max. Allowable Percent NcF (M) | 7.59%
17 | Acceptability Criterion (to accept, P<M) | Accepted

## Intermediate
ANSI Standard for Intermediate Flow
Test Results
1 | Sample Size: n | 3 | 100
2 | Sum of Measurements | 302 | 101
3 | Sum of Squared Measurements | 30402 | 101
4 | Correction Factor (CF) | 30401.3333
5 | Corrected Sum of Squares (SS) | 0.66667
6 | Variance (V) | 0.33333
7 | Estimate of Lot Standard Deviation | 0.57735
8 | Sample Mean | 100.6667
9 | Upper Specification Limit | 101.5
10 | Lower Specification Limit | 98.5
11 | Quality Index: QU (upper) | 1.4434
12 | Quality Index: QL (lower) | 3.7528
ANSI Standard Table B-5 used to derive values below
13 | Est. of Lot Percent NcF above Upper | 0%
14 | Est. of Lot Percent NcF below Lower | 0%
15 | Total Est. Percent NcF in Lot (p) | 0%
16 | Max. Allowable Percent NcF (M) | 7.59%
17 | Acceptability Criterion (to accept, P<M) | Accepted

## Minimum
ANSI Standard for Minimum Flow
Test Results
1 | Sample Size: n | 3 | 98
2 | Sum of Measurements | 292 | 97
3 | Sum of Squared Measurements | 28422 | 97
4 | Correction Factor (CF) | 28421.3333
5 | Corrected Sum of Squares (SS) | 0.66667
6 | Variance (V) | 0.33333
7 | Estimate of Lot Standard Deviation | 0.57735
8 | Sample Mean | 97.3333
9 | Lower Specification Limit | 90
10 | Quality Index: QL (lower) | 12.7017
ANSI Standard Table B-5 used to derive values below
11 | Est. of Lot Percent NcF | 0%
12 | Max. Allowable Percent NcF (M) | 40.47%
13 | Acceptability Criterion (to accept, P<M) | Accepted